## Med.Equipment
S/N | Items Description | Quantity | Specifications
1 | CBC Machine | 3 | 21 parameters with P_LCC, P_LCR and RBC, WBC, PLT histograms Throughput up to 60 samples per hour with double chambers design lmpedance method for WBC, RBC and PLT Photoelectric colorimetry method for HGB with cyanide free reagent Micro sample volume at 10uL for whole blood mode and 20uL for pre-diluted mode On board storage capacity of 100,000 sample results Optional reagent close system with external bar code scanning or key in Four kinds of QC modes: L-J, X.X-R and X-B Support USB data backup & system upgrade Built-in thermal printer LIS and HIS available with HL7 protocol
2 | Oxygine Machine 5L | 1 | 1. Flow Range: 1~5L/min (7F-5); 2~8L/min (7F-8); 2~10L/min (7F-10) 2. Oxygen Concentration: 95.5%~87% 3. Maximum Outlet Pressure: 45kPa± 4.5kPa (6.52psi± 0.65psi) (7F-5) Maximum Outlet Pressure: 45kPa± 4.5kPa (6.52psi± 0.65psi) (7F-8) Maximum Outlet Pressure: 62kPa± 4.5kPa (8.99psi± 0.65psi) (7F-10) 4. Pressure Relief Mechanism Operational at: 250kPa± 25kPa (36.25psi± 3.63psi) 5. Altitude: Up to 1828 meters (6000 ft) above sea level without degradation of concentration levels. From 1828 meters (6000 ft) to 4000meters (13129 ft) below 90% efficiency. (7F-5) Altitude: Up to 1230 meters (4000 ft) above sea level without degradation of concentration levels. From 1230 meters (4000 ft) to 4000meters (13129 ft) below 90% efficiency. (7F-8, 7F-10) 6. Sound Level: ≤ 60dB(A) 7. Maximum nebulizing rate (apply in the oxygen concentrators having the function of nebulizing): ≥ 0.2 mL/min 8. Power Supply: AC120V± 10% AC220V± 10% AC230V± 10% 50Hz / 60Hz 4.6amps(AC120V); 2.2amps(AC220V~240V) (7F-5, 7F-8) 8.5amps(AC120V); 4amps(AC220V~240V) (7F-10) 9. Input Power: ≤ 600VA (7F-5, 7F-8) ≤ 1000VA (7F-10)
3 | Microcope | 2 | 45mm DIN achromatic objective:4x,10x,40x(s) ,100x(s,oil) wide field plane-scope eyepiece:WF10X,WF16X; Total magnification:40x-1600x; interpupillary distance:55-75mm;45°inclined,360°rotatable double layer mechanical stage: 142mmx132mm,muving range XY:75x50mm; dia2-30mm iris diaphragm and dia32 filter; ABBE condensor:N.A. 1.25; Coaxial coarse and fine focusing adjustable mechanism:25mm,precision:0.002mm; built-in adjustable brightness halogen lamp 6V/20W
4 | Centrifuge 6 holes | 2 | Plug Type: 110V US Plug, 220V EU Plug Model: 80-1 laboratory centrifuge Capacity: 20ml x 6 (Angle Type) Speed Range: Start -4000 rpm Motor Power: 25W Timing Range: 1-60 minutes
5 | Autoclave-Electric 18L | 3 | Capacity 18 Liters Material High-quality 304 stainless steel chamber and body Power 1.2 kW to 2.4 kW Voltage Typically 110V or 220V/240V, 50/60 Hz Working Temperature 121°C to 126°C (250°F to 260°F) Working Pressure 0.12 MPa to 0.16 MPa (approx. 15-23 PSI or 1.2-1.6 bar) Max. Safety Pressure Around 0.165 MPa Sterilization Time Adjustable, commonly 0-60 minutes, with preset cycles for different loads Safety Features Safety valve for overpressure, overheat protection, and mechanical/electrical door interlocks Display Analog pressure gauge or digital LCD display Configuration Portable/benchtop design
6 | Sunction Machine | 1 | The electrical suction device models 7A-23D and 7A-23B operate silently, durable. Zero maintenance oil-free piston pumps. White plastic boards used to make the body look very elegant and make it simple to clean. Both designs feature accurate vacuum adjustable knobs that are simple to read. Two 2.5 liter glass jars each provide sufficient ability for any type of operation. They come with self-locking rubber caps of high quality. A foot switch is supplied with each suction unit as a standing accessory. Even if energy is cut-off, the suction devices retain negative pressure. In the case of a power failure, this prevents re-gorging of liquids. Protection against overflow is given in both jars.
7 | Blood Pressure (BP) + Stethescope | 8 | sphygmomanometer + stethoscope Economic sphygmomanometers with nylon cuff with velcro closure,reliable manometer with clip to be fixed on the cuff. Scale 0-300 mmHg. Air release at closed tap max 4 mmHg/minute. Manual setting of deflation possible up to ? mmHg/sec. From 260 mmHg to 15 mmHg max deflation time 10 sec. Scale graduated every 2 mmHg. Floating zero (the pointer has no stop point but swings freely), nylon velcro cuff, latex bulb with chromium plated valve. Chromium-plated air tap with air release regulation by screw valve. Supplied with a simulated-leather soft pouch with zip in a carton box. Available in 5 versions: standard, with one piece cuff, latex free, with incorporated stethoscopes or kit (sphygmomanometer+stethoscope).
8 | Digital Thermometer | 24 | Key Specifications: Measurement Range: The lowest to highest temperatures the device can read (e.g., -58°F to 572°F / -50°C to 300°C). Accuracy: How close the reading is to the true temperature (e.g., ±1°F or ±0.1°C), often specified within certain temperature ranges. Resolution: The smallest temperature increment displayed (e.g., 0.1°C or 1°F). Display: Usually an LCD, sometimes with backlighting or high-contrast digits. Probe Type: For probe models, material (stainless steel) and tip size (fine needle for meats). Response Time: How quickly it updates, often 1-3 seconds for fast-response models, notes TFA Dostmann. Power: Battery type (e.g., CR2032, AAA) and life (hours of use). Features: °C/°F switch, Auto-Off, Hold function, Max/Min memory, waterproof design, audible alerts, magnet for storage, and laser targeting (infrared).
9 | Glucometer with Strips (Oncall plus) | 8 | Sample: Fresh capillary whole blood Sample volume: Minimum 1.5uL Assay method: Glucose oxidase biosensor Unit of measure:mmol/L (mg/dl optional) Strip: URIT G26 glucose test strip Test time: 5 seconds Measuring range: 1. 1mmol/L-33. 3mmol/L Display: LCD Memory: 250 test results Quality control: Control solution,control strip Power supply: One CR2032 lithium battery(DC 3V) Battery capacity: up to 1000 tests Function: Downloadable 250-value memory with 7-,14- and 30-day averagingAutomatic temperature monitorAutomatic battery monitor and low power alarmSupport error message Ambient Condition: Temperature: 5-45ºC Relative humidity: ≤80%
10 | Pulse Oximeter Hand-held | 8 | Design tested by manufacturer in accordance with ISO 80601-2-61. Measurement: SpO2 and pulse rate monitor, with plethysmography waveform, for adults, children and neonates. SpO2 detection includes the range 70 - 100 %. Also capable of measuring lower saturations. SpO2 resolution 1 %. SpO2 accuracy (in the range 70 - 100%) within ± 3%, for all Features: Audible and visual alarms for low/high saturation and pulse rate. Audible and visual alarms for sensor error or disconnected, system errors, low battery. Alarm override and temporary silencing function. Capable of working with adult, paediatric and neonatal reusable probes. Enclosure protection IPX1 or better. Suitable for cleaning and disinfection with hospital-grade cleaning products. Device and probe weight < 200 g. Electrical characteristics: Operated by replaceable battery power supply, either rechargeable or single use. External AC battery charger, plug style and voltage as per local supply, as specified by purchaser. Running time on battery only > 12 hours.
11 | Ambubag Pediatric | 8 | Bag volume.infant (635 ml), Dimensions. Adult (295x127 mm), Pediatric (234x99 mm), Neonate (165x70 mm) Weight. Adult (350 g), Pediatric (230 g), Neonate (112 g)
12 | Delivery lamp | 5 | Examination Lamp is intended for the use in minor surgery room, out patient examination and doctor’s office. Minor lamp head with diameter of 100 mm enables close use without blocking out the vision of users. The 700mm-long flexible arm can freely position in order to fully satisfy any requirements for diverse illuminated angles. Handle is provided for flexible movability. The glass color filter creates white and soft light as natural light. Available in plastic/ metallic base. Lamp head Diameter: 100 mm Spot Diameter: 80-150 mm Height adjustable : 700 mm Power supply : 220V/50 Hz Power input: 50VA Specification of Bulb : 50W/12V
13 | Delivery set (Complete set) | 5 | Key Components: For Attendant/Hygiene: Soap, sterile gloves (latex/nitrile), face mask, cap, shoe covers, gown, apron, antiseptic wipes/solution (Povidone-Iodine). For Delivery Surface: Plastic sheet/drawsheet, absorbent pads, perineal pads/sheets, baby receiving sheet/towel. For Umbilical Cord Care: Sterile clamps (2-3), sterile blade/scalpel, clean string/ties. For Newborn Care: Baby wrap/blanket, mucus extractor, baby wipes, feeding bottle (sometimes). For Postpartum: Sanitary pads, placenta bag (disposable).
14 | Wheelchair | 2 | Key Dimensions Seat Width: Varies (e.g., 16-22"+), measured across hips plus a couple of inches for comfort. Seat Depth: From buttocks to back of knee, often with a 2-4 finger gap. Seat Height: From seat to floor (around 19.5" is common). Overall Width/Length: Affects maneuverability and doorway clearance (e.g., 25" width, 42" length typical). Backrest Height: For support (e.g., 16" for active, 20"+ for high support). Common Components & Features Frame: Folding for portability or rigid for stability, often lightweight. Wheels: 24" rear wheels for propulsion, smaller front castors (e.g., 8"). Armrests: Fixed, flip-up, or removable, sometimes desk-length. Footrests: Swing-away for easier transfers, flip-up plates. Brakes: Locking on rear wheels. Upholstery: Canvas or easy-clean materials. Important Considerations Weight Capacity: Varies by model (e.g., 100-115kg / 220-250lbs typical). Adjustability: Seat height, depth, and footrest positions can often be adjusted. Maneuverability: Affected by wheel size, frame, and camber (rear wheel angle).
15 | Examination bed | 2 | Dimensions: Common sizes are approximately 190-210 cm (Length) x 60-66 cm (Width) x 66-80 cm (Height), with variations for multi-section or specialty tables. Adjustability: Height: Often adjustable (manual or electric) for patient transfer and clinician access (e.g., 470mm to 920mm). Sections: Typically 2 or 3 sections (head, back, seat) with adjustable angles (e.g., headrest -45° to +40°). Load Capacity: Minimum 150 kg, with higher-end models supporting 225 kg or more. Materials: Frame: CRCA steel, powder-coated steel, or stainless steel for durability. Upholstery: Fire-retardant, waterproof, anti-static PVC or similar synthetic leather with dense foam (e.g., 4-6 cm thick)
16 | Scissors | 8 | Material: High-grade Stainless Steel (often German); sometimes with Tungsten Carbide (TC) cutting edges for added hardness and longevity. Design & Geometry: Curvature: Straight (for superficial cuts) or Curved (for deeper dissection). Tip: Blunt/Blunt (for materials like gauze), Sharp/Sharp (tough tissue), or Sharp/Blunt (general purpose). Length: Ranges from small (e.g., 3.5" Iris) to large (10"+ Mayo), influencing application. Shank-to-Blade Ratio: Longer shanks for deeper reach. Finish: Fine finish, satin (low reflective), or polished. Application: General Surgery, Dermatology, Emergency, Bandage Cutting (Lister), Delicate Dissection (Metzenbaum). Handling: Finger rings (sometimes oversized), balanced weight, ergonomic design for control. Durability: Rustproof, autoclavable, resistant to corrosion.
17 | Privacy Screen | 1 | Key Specifications Material: Frames are typically made of stainless steel or powder-coated iron/aluminum tubing for durability, corrosion resistance, and a smooth, easy-to-clean surface. The panels or curtains are often a heavy-duty, opaque, and fluid-resistant fabric like nylon or polyester-cotton blends. Hygiene & Cleaning: Materials must be non-porous and easy to sanitize with hospital-grade disinfectants, including alcohol and bleach solutions, to prevent the spread of infection. Some advanced options offer anti-bacterial materials or high-pressure laminate finishes. The fabric panels are often detachable for easy washing. Mobility & Storage: Most hospital partitions are mobile, featuring silent, 366-degree swivel casters (wheels) for easy movement. These wheels typically have brakes for secure positioning. The panels are connected by durable hinges, allowing them to be folded into a compact size (concertina design) for storage when not in use. Dimensions: Standard heights are around 180 cm (approx. 6 feet or 70.9 inches) to ensure full visual privacy. Individual panel widths are often about 50-60 cm (approx. 19.7-24 inches), with overall widths varying based on the number of panels (e.g., 3-panel, 4-panel, or 6-panel options). Safety & Design: Opaque: The fabric must be completely opaque to ensure patient dignity and privacy. Airflow: Many models feature a mesh top section to allow for air circulation and proper function of ceiling sprinkler systems. Stability: Low-profile feet or a sturdy frame design help prevent tripping hazards and ensure stability, whether the screen is used in a straight line or an angled configuration.
18 | Baby Scale | 3 | Capacity: Typically up to 15kg or 20kg. Resolution/Graduation: 5g or 10g for precise readings. Accuracy: ±5g (often required). Display: Large, clear LCD. Functions: TARE: Zero out the weight of blankets/towels. HOLD/Auto-HOLD: Locks the reading for stability. Auto-off: Saves battery. Dual-Range: Switches from fine (5g) to coarse (10g) for toddlers. BMIF (Breast-Milk-Intake-Function): Measures milk consumed. Material: Smooth, non-absorbent ABS plastic for easy cleaning. Power: Battery (AA) or AC adapter. Certifications: Often CE or MDD Class III approved for medical use. Design: Lightweight, portable, splash-proof, with a smooth, curved tray.
19 | IV stand | 4 | Key Specifications Height Adjustment: Telescoping poles allow range, e.g., 1350mm to 2150mm (OR) or 1520mm (ICU). Material: Often stainless steel (like 304 SS) for durability, corrosion resistance, and hygiene; sometimes lightweight aluminum. Base: Legs: 5-legged (penta-base) for stability, sometimes with castors (wheels). Weight: Weighted bases offer stability, especially for heavy loads. Poles: Inner tube diameter around 18mm, outer around 25mm. Hooks/Hangars: Typically 2 to 4 hooks, sometimes more, for hanging bags/pumps. Load Capacity: Varies by model, e.g., up to 10kg (portable) or 15kg (heavy-duty). Finish: Polished or powder-coated for easy cleaning and infection control.
TOTAL